## Slide 1: Constructing spherical panoramas of a bladder phantom from endoscopic video using bundle adjustment
SPIE Medical Imaging 2011,
Visualization, Image-Guided Procedures, and Modeling
Tuesday February 15th, 2011
Timothy D. Soper, John E. Chandler, Michael P. Porter, Eric J. Seibel

## Slide 2: Bladder Surveillance
www.nhs.uk
Bladder Cancer:
  5th most prevalent in U.S.1
  70,530 new cases yearly
  Recurrence 50%
Flexible Cystoscopy
  Annual, semi-annual, or quarterly surveillance
  ~5mm diameter
  Local anesthesia
Rigid Cystoscopy
  Biopsy or resection
  General anesthesia
1 Jemal, A., Siegel, R., Xu, J. et al., “Cancer statistics, 2010,” CA Cancer J Clin, 60(5), 277-300 (2010).

## Slide 3: A New Approach to Bladder Surveillance
Ultrathin (1.5 mm)
Highly flexible
High Resolution (500 x 500)
Full-Color
30 Hz video frame rate
Scanning Fiber Endoscope
Advancement
Patient comfort
Laser diagnostics/therapies
Automation
Replace urologist with nurse/technician
Implement automated scanning device
Urologist performs post-procedural review
Proposed Model
Yoon, Sangtae, Reinhall, et al., 2009
Panorama

## Slide 4: Image Alignment
Requirements for full panoramic
reconstruction :
Nonplanar
Free moving endoscope
360° field of view
Global alignment
Self calibrating
Not real time
Challenges:
Presumed planar geometry
Global alignment
Precalibration

## Slide 5: Structure from Motion
Simultaneous recovery of scene geometry and camera motion
Virtual Tourism1
1Snavely, N., Seitz, S. M., and Szeliski, R., ACM Transactions on Graphics, 2006.
[R, t ]

## Slide 6: Software
Frame selection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO

## Slide 7: Frame Selection
Frame selection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
1Lowe, International Journal of Computer Vision, 2004
Scale Invariant Feature Transform1
Feature matching
Random sample concensus (RANSAC)
Sparse selection
1
N

## Slide 8: Frame Matching
+3
+3
Associated matching
Nonsequential matching
Sequential matching
Frame selection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO

## Slide 9: Track Assignment
Frame selection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
For all features consistently matched between 3+ frames:
assign as track

## Slide 10: Bundle Adjustment
Non-linear least squares minimization between the observed and predicted pixel positions:
Frame selection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
Large parameter sets solved by exploiting sparsity!
)
Camera Equations:
(3D rigid body)
(projective)
(radial distortion)
(camera
equations)
State variable:
Update:
Levenberg-Marquardt
3D point parameters and camera parameters do not combine
Each camera sees only a subset of points

## Slide 11: Bundle Adjustment
Frame selection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
Implementation:
Constrain mosaic to spherical surface
Constrain frames to lie within sphere.
Incremental adjustment
add and initialize first frame
add visible tracks and project
bundle adjust
add frames and estimate pose
bundle adjust

## Slide 12: Experiment 1: Bladder Phantom
240 frames
Projection error = 0.47 pixels

## Slide 13: Experiment 2: Excised Pig Bladder
857 frames
Projection error = 1.41 pixels

## Slide 14: Experiment 2: Reconstruction 2D

## Slide 15: Conclusions
Several hundred frames aligned with pixel level accuracy.
No calibration necessary
Spherical model proved adequate
Sufficient overlap is essential!

## Slide 16: Future Work
Validate automated scanning mechanism
Improve blending, deghosting
Develop interface

## Slide 17: Acknowledgements
Rich Johnston
Charles Melville
Cameron Lee
Coulter Foundation & University of Washington Department of Bioengineering
University of Washington Center for Commercialization
email: tsoper@uw.edu
Human Photonics Laboratory
Funding

## Slide 19
Animation by Mr. Duff Hendrickson, Seattle, WA, copyright University of Washington.
Scanning Fiber Endoscope (SFE)

## Slide 20: Full Panoramic Reconstruction
Simultaneously recover scene geometry and camera motional
Reconstruction Metrics:
Single mosaic of entire bladder
No presumed shape
Global alignment
No Calibration
Structure from Motion:
Virtual Tourism1
N
1Snavely, N., Seitz, S. M., and Szeliski, R., ACM Transactions on Graphics, 2006.

## Slide 21: Bundle Adjustment
Non-linear least squares minimization between the observed and predicted pixel positions:
Feature detection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
Large parameter sets solved by exploiting sparsity!
)
Camera Equations:
(3D rigid body)
(projective)
(radial distortion)
(camera
equations)
State variable:
Levenberg-Marquardt
3D point parameters and camera parameters do not combine
Each camera sees only a subset of points

## Slide 22: Constructing Planar Panoramas of the Bladder
Planar Homography:
Limitations:
Assumes planar scene
  misalignment
  pixel stretching
Limited FOV
Lacks global alignment
Requires precalibration
H

## Slide 23: Frame Matching
Feature detection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
+3
+3
Sequential frame matching

## Slide 24: Frame Matching
Feature detection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
+3
+3
Nonsequential frame matching

## Slide 25: Frame Matching
Frame selection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
Feature Matching (step1):
Match each point to it’s nearest neighbor in feature space
Feature Matching (step2):
RANdom SAmple Consensus (RANSAC):
Compute homography (H )
with the most inliers

## Slide 26: Feature Detection
Frame selection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
Scale Invariant Feature Transform (SIFT1)
Scale, rotation invariant
Keypoints
Descriptors
1Lowe, International Journal of Computer Vision, 2004

## Slide 27: Frame Matching
Feature detection
Frame matching
Track assignment
Bundle adjustment
Stitching
VIDEO
+3
+3
Associated frame matching