Tektronix 2400 Series Digital Oscilloscopes

_____________________________________________________________________________

Introduction:

This  instrument  module  provides  GPIB  programming  support  for  the

Tektronix 2400 series of digital oscilloscope - 2430A, 2432, and 2440.  It

contains functions for opening, configuring, reading  waveforms  from, and

manipulating the cursor capabilities of the instrument.

_____________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

- one of the indicated scopes is connected to the GPIB.

- the  GPIB address supplied to the initialize function must match

the current GPIB address of the scope.

- there is one unused device name in the handler.

_____________________________________________________________________________

LabWindows Libraries:

In  order to use this instrument module outside of the LabWindows envi-

ronment, it must be linked with the following LabWindows libraries:

- Formatting and I/O

- GPIB

- Analysis

_____________________________________________________________________________

Error Codes:

All  error codes are  returned  in a global integer variable whose name

depends on the language being used:

For QuickBASIC : tek2400.err%

For C          : tek2400_err

A program should examine this variable after each function call is made

to determine if any errors occurred.

_____________________________________________________________________________

Function Tree Layout:

Initialize

Configure

Vertical

Horizontal

A Trigger

B Trigger

Auto Setup

Acquisition Settings

Save Settings

Read Waveform

Save to Reference

Measurement Functions

Setup Measurement

Read Measurement

Cursor Functions

Set Cursor

Read Cursor

Close

_____________________________________________________________________________

Function           : Initialize

QuickBASIC Syntax  : CALL tek2400.init(Addr%)

C Syntax           : void tek2400_init(Addr);

int Addr;

Purpose            : Initializes the 2340A on the GPIB

Parameter List     : Addr - GPIB Address (0 to 30)

QuickBASIC Example : CALL tek2400.init(1)

C Example          : tek2400_init(1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : GPIB Address out of range

220 : Unable to open instrument

223 : Device at specified address is not a 2400

series oscilloscope

230 : GPIB write error

231 : GPIB read error

233 : Unable to configure GPIB address

_____________________________________________________________________________

Function           : Set Vertical Parameters

QuickBASIC Syntax  : CALL tek2400.set.vertical (Channel%, VoltsDiv#, Cou%,_

Inv%, Fif%)

C Syntax           : tek2400_set_vertical (Channel, VoltsDiv, Cou, Inv, Fif);

int Channel;

double VoltsDiv ;

int Cou;

int Inv;

int Fif ;

Purpose            : Configures the vertical volts/division, input coupling,

invert setting, and input impedance of the scope.

Parameter List     : Channel - channel of scope to configure

1 or 2

VoltDiv - volts per division setting

0.002 to 50.0, in 1-2-5 sequence

Cou - input coupling mode

0 : AC

1 : DC

2 : GND

Inv - invert off or on

0 : Off

1 : On

Fif - fifty-ohm input impedance off or on

0 : Off

1 : On

QuickBASIC Example : CALL tek2400.set.vertical(1, 10.0, 0, 0, 0)

C Example          : tek2400_set_vertical(1, 10.0, 0, 0, 0);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Channel out of range

-2 : Volts/Div out of range

-3 : Coupling out of range

-4 : Invert out of range

-5 : Fifty-ohm Setting out of range

230 : GPIB write error

231 : GPIB read error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set Vertical Mode

QuickBASIC Syntax  : CALL tek2400.set.vertical.mode (VertMode%)

C Example          : tek2400_set_vertical_mode (VertMode);

int VertMode ;

Purpose            : Configures the vertical display mode of the scope.

Parameter List     : VertMode - the vertical display mode

0 : Channel 1 only

1 : Channel 2 only

2 : Channel 1 and 2

3 : Channel 1 Versus Channel 2 (XY)

QuickBASIC Example : CALL tek2400.set.vertical.mode(1)

C Example          : tek2400_set_vertical_mode(1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Vertical Mode out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set Multiply Mode

QuickBASIC Syntax  : CALL tek2400.set.multiply (Mult%)

C Example          : tek2400_set_multiply (Mult);

int Mult ;

Purpose            : Turns multiply mode off or on.

Parameter List     : Mult - multiply mode off or on

0 : Off

1 : On

QuickBASIC Example : CALL tek2400.set.multiply(1)

C Example          : tek2400_set_multiply(1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Multiply out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set Add Mode

QuickBASIC Syntax  : CALL tek2400.set.add (Add%)

C Example          : tek2400_set_add (Add);

int Add ;

Purpose            : Turns add mode off or on.

Parameter List     : Add - add mode off or on

0 : Off

1 : On

QuickBASIC Example : CALL tek2400.set.add(1)

C Example          : tek2400_set_add(1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Add out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set Bandwidth

QuickBASIC Syntax  : CALL tek2400.set.bw (BW)

C Example          : tek2400_set_bw (BW);

int BW ;

Purpose            : Sets the bandwidth of the scope.

Parameter List     : BW - the desired bandwidth setting

0 : Full

1 : 50 MHz for 2430A/2432

100 MHz for 2440

2 : 20 MHz

QuickBASIC Example : CALL tek2400.set.bw(1)

C Example          : tek2400_set_bw(1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Bandwidth out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set Horizontal Parameters

QuickBASIC Syntax  : CALL tek2400.set.horizontal (HorMode%, ASecDiv#, _

BSecDiv#, Delay#, Pos%)

C Syntax           : tek2400_set_horizontal (HorMode, ASecDiv, BSecDiv,

Delay, Pos)

int HorMode ;

double ASecDiv ;

double BSecDiv ;

double Delay ;

int Pos ;

Purpose            : Configures the horizontal mode, sweep time for time-

base A and B, the delay, and position

Parameter List     : HorMode - the horizontal mode

0 : A mode

1 : A with B intensified

2 : B mode

ASecDiv - seconds per division for timebase A

2430A : 5 nS to 5 s

2432, 2440 : 2 nS to 5 s

BSecDiv - seconds per division for timebase B

All scopes : 5 nS to 5 s

Delay - delay time used by B timebase relative to A

dependent upon scope and BSecDiv; refer to the

user manual for your scope

Pos - horizontal position of traces on the display

0 to 1023

QuickBASIC Example : CALL tek2400.set.horzontal(1,10.0e-6, 100.e-6, 1.0e-3,_

512)

C Example          : tek2400_set_horizontal(1, 10.0e-6, 100.e-6, 1.0e-3,512);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Vertical Mode out of range

-2 : A Sec/Div out of range

-3 : B Sec/Div out of range

-4 : Delay out of range

-5 : Position out of range

230 : GPIB write error

231 : GPIB read error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set A Trigger

QuickBASIC Syntax  : CALL tek2400.set.a.trigger (Level#, Slope%, TPos%,_

Mode%, Source%, Cou%)

C Syntax           : tek2400_set_a_trigger (Level, Slope, TPos, Mode,

Source, Cou);

double Level ;

int Slope ;

int TPos ;

int Mode ;

int Source ;

int Cou ;

Purpose            : Configures the A trigger settings of the scope.

Parameter List     : Level - trigger level in volts

Slope - determines whether trigger occurs on rising

(positive) or falling (negative) transition of trigger

signal

0 : Negative

1 : Positive

TPos - number of data points acquired prior to trigger

with (1023 -[Tpos * 32]) points acquired after trigger

1 to 30

Mode - trigger mode for trigger A

0 : Auto level

1 : Auto

2 : Normal

3 : Single

Source - source for A trigger signal

0 : Vertical

1 : Ch 1

2 : Ch 2

3 : External 1

4 : External 2

5 : Line

6 : Word

Cou - input coupling for A trigger signal input

0 : DC

1 : AC

2 : Noise reject

3 : HF reject

4 : LF reject

QuickBASIC Example : CALL tek2400.set.a.trigger (2.0, 0, 31, 0, 1, 1)

C Example          : tek2400_set_a_trigger (2.0, 0, 31, 0, 1, 1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Level out of range

-2 : Slope out of range

-3 : Trigger Position out of range

-4 : Mode out of range

-5 : Source out of range

-6 : Coupling out of range

230 : GPIB write error

231 : GPIB read error

232 : Instrument not initialized

_____________________________________________________________________________

Control Name:  Ext1 / 5

Variable Type: Integer

Valid Values:  0   Do not divide by 5

_____________________________________________________________________________

Function           : Set Gain of External input 1

QuickBASIC Syntax  : CALL tek2400.set.ext1.div5 (Div5%)

C Example          : tek2400_set_ext1_div5 (Div5);

int Div5;

Purpose            : Indicates whether or not to divide the signal from

External input 1 by a factor of 5.

Parameter List     : Div5

0 : Do not divide by 5

1 : Divide by 5

QuickBASIC Example : CALL tek2400.set.ext1.div5(1)

C Example          : tek2400_set_ext1_div5(1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Div5 out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set Gain of External input 2

Purpose            : Indicates whether or not to divide the signal from

External input 2 by a factor of 5.

QuickBASIC Syntax  : CALL tek2400.set.ext2.div5 (Div5%)

C Example          : tek2400_set_ext2_div5 (Div5);

int Div5;

Parameter List     : Div5

0 : Do not divide by 5

1 : Divide by 5

QuickBASIC Example : CALL tek2400.set.ext2.div5(1)

C Example          : tek2400_set_ext2_div5(1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Div5 out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set B Trigger

QuickBASIC Syntax  : CALL tek2400.set.b.trigger (Level#, Slope%, TPos%, _

Mode%, Source%, Cou%)

C Syntax           : tek2400_set_a_trigger (Level, Slope, TPos, Mode,

Source, Cou);

double Level ;

int Slope ;

int TPos ;

int Mode ;

int Source ;

int Cou ;

Purpose            : Configures the B trigger settings of the scope.

Parameter List     : Level - trigger level in volts

Slope - determines whether trigger occurs on rising

(positive) or falling (negative) transition of trigger

signal

0 : Negative

1 : Positive

TPos - number of data points acquired prior to trigger

with (1023 -[Tpos * 32]) points acquired after trigger

1 to 30

Mode - selects whether B sweep free-runs or must be

triggered to run

0 : Free run

1 : Trigger required

Source - source for B trigger signal

0 : Vertical

1 : Ch 1

2 : Ch 2

3 : External 1

4 : External 2

5 : Word

Cou - input coupling for B trigger signal input

0 : DC

1 : AC

2 : Noise reject

3 : HF reject

4 : LF reject

QuickBASIC Example : CALL tek2400.set.b.trigger (2.0, 0, 31, 0, 1, 1)

C Example          : tek2400_set_b_trigger (2.0, 0, 31, 0, 1, 1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Level out of range

-2 : Slope out of range

-3 : Trigger Position out of range

-4 : Mode out of range

-5 : Source out of range

-6 : Coupling out of range

230 : GPIB write error

231 : GPIB read error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Auto Setup

QuickBASIC Syntax  : CALL tek2400.auto.setup (Mode%, Resolution%)

C Syntax           : tek2400_auto_setup (Mode, Resolution)

int Mode ;

int Resolution ;

Purpose            : Causes scope to automatically adjust vertical, hori-

zontal, and trigger parameters to display the input

waveform

Parameter List     : Mode - specifies the characteristic of the input signal

to use for optimizing the setup of the scope

0 : View

1 : Period

2 : Rise

3 : Fall

4 : Pulse

Resolution - when set to HI, causes waveform to be

spread over entire 20 division acquisition window;

LO results in optimization of horizontal resolution

to fit the waveform on the scope screen

0 : HI

1 : LO

QuickBASIC Example :   CALL tek2400.auto.setup(0, 1)

C Example          :   tek2400_auto_setup(0, 1);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Mode out of range

-2 : Resolution out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Set Acquisition Parameters

QuickBASIC Syntax  : CALL tek2400.acq.set (AcqMode%, Delta%, Rep%, Env%, _

Avg%, Ref1%, Ref2%,Ref3%, Ref4%, RunMode%)

C Syntax           : tek2400_acq_set (AcqMode, Delta, Rep, Env, Avg, Ref1,

Ref2, Ref3, Ref4, RunMode);

int AcqMode;

int Delta;

int Rep;

int Env;

int Avg;

int Ref1;

int Ref2;

int Ref3;

int Ref4;

int RunMode;

Parameter List     : AcqMode - sets the acquisition mode to normal, enve-

lope, or average

0 : Normal

1 : Envelope

2 : Average

Delta - turns Save-On-Delta mode off or on

0 : Save-On-Delta off

1 : Save-On-Delta on

Rep - turns repetitive sampling mode off or on

0 : Repetitive mode off

1 : Repetitive mode on

Env - sets the number of acquisitions over which

which to accumalate envlopes (maximum and

minimum values).

0 : continuous

1 to 256, powers of 2 : number of envelopes

Avg - number of sweeps to average before starting over

2 to 256, powers of 2

Ref1 - specifies whether or not Reference 1 is

displayed

0 : Do not display

1 : Display

Ref2 - specifies whether or not Reference 2 is

displayed

0 : Do not display

1 : Display

Ref3 - specifies whether or not Reference 3 is

displayed

0 : Do not display

1 : Display

Ref4 - specifies whether or not Reference 4 is

displayed

0 : Do not display

1 : Display

RunMode - sets the action to occur when the function

is executed

0 : Leave current setting

1 : Start an acquisition

2 : Change to Save Mode

QuickBASIC Example :   CALL tek2400.acq.set (0, 0, 1, 1, 2, 0, 0, 0, 0, 0)

C Example          :   tek2400_acq_set (0, 0, 1, 1, 2, 0, 0, 0, 0, 0);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Acquisition Mode out of range

-2 : Save-on-Delta out of range

-3 : Repetitive Mode out of range

-4 : Envelopes out of range

-5 : Averages out of range

-6 : Ref1 out of range

-7 : Ref2 out of range

-8 : Ref3 out of range

-9 : Ref4 out of range

-10 : Run Mode out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Save Settings

QuickBASIC Syntax  : CALL tek2400.save.set(Filename$, Read/write%)

C Syntax           : void tek2400_save_set(Filename, Read/Write);

char *Filename;

int Read/Write;

Purpose            : Reads acquisition settings from a file and sends them

to the scope or writes the current acquisition

settings of the scope to a file depending on the

Read/Write mode.

Parameter List     : Filename - specifies the name of the file from which

to read or write.

Read/Write - specifies whether to read settings from

file or to write current settings to file.

QuickBasic Example : CALL tek2400.save.set ("tek2400.set", 0)

C Example          : tek2400_save_set ("tek2400.set", 0);

Error Codes        : The valid error codes for this function are:

-1 : Read/Write out of range

-2 : Error reading or writing to file

-230 : GPIB write error

-231 : GPIB read error

-232 : Instrument is not initialized

400 : Error in opening or closing a file

401 : Error in reading from or writing to

a file

_____________________________________________________________________________

Function           : Read Waveform

QuickBASIC Syntax  : CALL tek2400.read.wfm(Source%,Wfm#(),SamplePer#, TrigPos#)

C Syntax           : void tek2400_read_wfm(Source,Wfm,SamplePer,TrigPos);

int Source;

double Wfm[1024];

double *SamplePer ;

double *TrigPos ;

Purpose            : Reads a waveform from an acquisition source or from a

reference location.

Parameter List     : Source - specifies the source of the waveform

1 : Acquire from Channel 1

2 : Acquire from Channel 2

3 : Acquire waveform from Add operation

4 : Acquire waveform from Multiply operation

5 : Read Ref 1

6 : Read Ref 2

7 : Read Ref 3

8 : Read Ref 4

Wfm - a real array variable in which the wave-

form data  is returned. This  array must

be at least 1024 elements long.

SamplePer - a real variable in which the value

of the sampling period used by the

oscilloscope to digitize the wave-

form is returned.

TrigPos - a real variable indicating the position of

the trigger point in the waveform, i.e. the

sample number representing time 0.

QuickBASIC Example : DIM wfm.data#(1024)

CALL tek2400.read.wfm(1,wfm.data#(),period#, t.zero#)

C Example          : double period;

double t_zero ;

double wfm_data[1024];

tek2400_read_wfm(1,wfm_data,&period, &t_zero);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Source out of range

230 : GPIB write error

231 : GPIB read error

232 : Instrument not initialized

300 : Specified waveform source is not valid for

current instrument configuration

_____________________________________________________________________________

Function           : Save to Reference

QuickBASIC Syntax  : CALL tek2400.save.to.reference(Source%,Destination%)

C Syntax           : void tek2400_save_to_reference(Source,Destination);

int Source;

int Destination ;

Purpose            : Saves a waveform to a reference location.

Parameter List     : Source - specifies the source of the waveform

1 : Acquire from Channel 1

2 : Acquire from Channel 2

3 : Acquire waveform from Add channel

4 : Acquire waveform from Multiply channel

5 : Ref 1

6 : Ref 2

7 : Ref 3

8 : Ref 4

Destination- specifies the location in which to save

the waveform

1 : Ref 1

2 : Ref 2

3 : Ref 3

4 : Ref 4

QuickBASIC Example : CALL tek2400.save.to.reference(1, 4)

C Example          : tek2400_save_to_reference(1, 4);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Source out of range

-2 : Destination out of range

230 : GPIB write error

231 : GPIB read error

232 : Instrument not initialized

300 : Specified waveform source is not valid for

current instrument configuration

_____________________________________________________________________________

Function           : Setup Measurement

QuickBASIC Syntax  : CALL tek2400.setup.measurement(Window%, Method%, Mark%)

C Syntax           : tek2400_setup_measurement (Window, Method, Mark);

int Window ;

int Method ;

int Mark ;

Purpose            : Configures the parameters for taking measurements.

Parameter List     : Window - turns measurement window off or on

0 : window off

1 : window on

Method - selects the method for determining Top and

Bottom of the waveform for measurements

0 : Min/Max

1 : Histogram

2 : Cursor (uses the current Volts cursor values)

Mark - turns the display threshold crossing marks on

or off

0 : marks off

1 : marks on

0 : Volts

QuickBASIC Example : CALL tek2400.setup.measurement (0, 1, 0)

C Example          : tek2400_setup_measurement (0, 1, 0);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Window out of range

-2 : Method out of range

-3 : Mark out of range

230 : GPIB write error

232 : Instrument not initialized

_____________________________________________________________________________

Function           : Read Measurement

QuickBASIC Syntax  : CALL tek2400.read.measurement(Source%, Type%, Reading#)

C Syntax           : tek2400_read_measurement (Source, Type, Reading);

int Source ;

int Type ;

double *Reading ;

Purpose            : Returns a single measurement.

Parameter List     : Source - specifies the location of the waveform on

which a measurement is made

1 : Chan 1

2 : Chan 2

3 : Add channel

4 : Multiply channel

Type - specifies the type of measurement to take

0 : Distal

1 : Mesial

2 : Proximal

3 : Max

4 : Mid

5 : Min

6 : Peak-to-Peak

7 : Top

8 : Base

9 : Mean

10 : Overshoot

11 : Undershoot

12 : RMS

13 : Area

14 : Width

15 : Duty

16 : Frequency

17 : Period

18 : Rise time

19 : Fall time

20 : Delay

Reading - real variable to contain the measurement

QuickBASIC Example : CALL tek2400.read.measurement (1, 0, reading#)

C Example          : double reading ;

tek2400_read_measurement (1, 0, &reading);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Source out of range

-2 : Type out of range

230 : GPIB write error

231 : GPIB read error

232 : Instrument not initialized

300 : Specified waveform source is not valid for

current instrument configuration

_____________________________________________________________________________

Function           : Set Cursor

QuickBASIC Syntax  : CALL tek2400.set.cursor(OnOff%,Type%,Units%,Mode%)

C Syntax           : void tek2400_set_cursor(OnOff,Type,Units,Mode);

int OnOff;

int Type;

int Units;

int Mode;

Purpose            : Configures the cursor parameters to the specified

values  and then  places the  instrument in local

mode.

Parameter List     : OnOff - cursors on/off mode

0 : cursors off

1 : cursors on

Type - measurement type

0 : Volts

1 : Time

2 : 1/Time

3 : Slope

Units - measurement units

0 : Volts, Seconds, Hertz, or Volts/Second

1 : Percent

2 : dB or Degrees

Mode - measurement mode

0 : Absolute

1 : Delta

QuickBASIC Example : CALL tek2400.set.cursor(1,0,0,0);

C Example          : tek2400_set_cursor(1,0,0,0);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : On/Off out of range

-2 : Type out of range

-3 : Units out of range

-4 : Mode out of range

230 : GPIB write error

232 : Instrument not initialized

234 : Unable to place instrument in local mode

_____________________________________________________________________________

Function           : Read Cursor

QuickBASIC Syntax  : CALL tek2400.read.cursor(Value#)

C Syntax           : void tek2400_read_cursor(Value);

double *Value;

Purpose            : Returns the current cursor measurement.

Parameter List     : Value - a real variable in which the value of

the cursor measurement is returned.

QuickBASIC Example : CALL tek2400.read.cursor(slope#);

C Example          : double slope;

tek2400_read_cursor(&slope);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : Cursors are turned off

230 : GPIB write error

231 : GPIB read error

232 : Instrument not initialized

234 : Unable to place instrument in local mode

_____________________________________________________________________________

Function           : Close

QuickBASIC Syntax  : CALL tek2400.close()

C Syntax           : void tek2400_close();

Purpose            : Effectively disconnects the 2340A from the GPIB

and returns the instrument to local mode.

Parameter List     : NONE

QuickBASIC Example : CALL tek2400.close()

C Example          : tek2400_close();

Error Codes        : The valid error codes for this function are:

0 : Success

221 : Unable to close instrument

232 : Instrument not initialized

_____________________________________________________________________________