#General Program Description

Included on this disk are the following spectral analysis programs:

GAMMA.COM, ENRAD.COM, ALPHA.COM, and XRAY.COM.

GAMMA and ENRAD are programs to display gamma emission spectral data.

ALPHA is intended to display alpha emission spectral data.  XRAY can display

x-ray diffraction pattern spectral data.  Sample spectra are included on this

disk.  These programs will display spectral data on the hi-resolution color

graphics screen and allow the user to examine it.  He/she can expand or

contract the spectrum and print it by using a print screen utility.  Selected

areas of the spectrum can be integrated, and peaks in a gamma or alpha

spectrum can be identified by comparison with an isotope library.  Any program

is started by typing the following command to the DOS:

#       command [filespec]

where #command# is either GAMMA, ENRAD, ALPHA, or XRAY; and where #[filespec]# is

a spectral file.  If the filespec is not included on the command line, then

the program will prompt the user to supply one.  Drive and path may be

included.  The spectral file must be in a particular format in order for the

program to properly read it (see section on file format).  The program then

reads the spectral file if it exists; if not, it will ask the user to supply

an existing filespec.

#Program Operation

Let's explore the operation of one of the spectral analysis programs by

looking at one of the sample spectra provided on the disk.  Type the following

command to get started:

#ENRAD corew4.ccs

We will start with the ENRAD program (ENRAD stands for environmental

radiation).  Corew4.ccs is a gamma spectrum of a sediment core taken from the

bottom of the Calcasieu Lake in southwest Louisiana.  We expect to see only

naturally occurring gamma emitters, so we choose to use the ENRAD program.

ENRAD will read this spectral file.

Next, the program will read in the isotope library file. (At this time,

XRAY does not include any kind of library file.)  This file must be present on

the default drive.  These files contain isotopic information which the program

uses to identify peaks in the gamma or alpha spectra.  (See section on Isotope

Libraries for a description of these files.)

Next, the program searches for a file containing the detector geometry

coefficients.  This file is called COEFFS.TXT.  It currently contains five

geometries (see section on coefficients file).

Next, the program will search the default drive for a file which has the

same filename and an extension of .CFG.  This file contains answers to the

questions asked by the program concerning whether or not to connect the dots

making up the spectral plot, and vertical scaling of the data.  If this file

is not present, the program prompts the user for this information.  The

questions asked are:

"Connect dots? (Y/N)".  Type Y or N in response to this question.

"Logarithmic plot? (Y/N)".  If the answer is Y, the program will scale

the data and display the spectrum.  If the answer is N, the program will ask

the user to input a scaling factor.  If the user just presses RETURN or enters

a zero, then the data will be scaled so that the largest value in the spectrum

will appear full scale on the graphic plot.  Choosing a smaller value will

result in expansion of the y�axis.  After the data is scaled, the spectrum is

displayed on the graphic screen and the user can begin to examine it.

We should see the dots not connected and a logarithmic plot.  The

vertical line near the middle of the screen is the graphic cursor.  It can be

moved by pressing the left and right arrow keys.  It can be moved faster by

holding down the control key and pressing the arrow keys.  As the cursor is

moved, its position is constantly updated on a status line near the bottom of

the screen.  The cursor's channel number is shown.  This is its position

within the spectral file.  For this particular spectrum, there are 1024

channels.  The energy corresponding to this channel is calculated and shown

(in units of kilo electron volts) if this spectrum has been calibrated.  This

calibration data is stored along with the spectrum.  Finally, the counts in

this channel are shown.

The graphic screen is a window through which we are looking at the

spectrum.  To move the window to the right, press the PgUp key.  To move the

window to the left, press the PgDn key.  Look at the help screen by pressing

the / or ? key.  When you are ready to quit the program, press control-Q.  The

program will return to the DOS.  A description of the function keys is given

below.

F1 LEFT Pressing the F1 key will place a marker on the left side of a

region of interest.

F2 RIGHT Pressing the F2 key will place a marker on the right side of a

region of interest.

(Any time both LEFT and RIGHT marks are placed on a region of interest,

this spectral region will be integrated and the results shown on a status line

near the bottom of the screen.  Shown are the LEFT and RIGHT channels, the

gross area and the net area.  The net area is calculated by drawing a baseline

from the left to the right mark, and subtracting the area below the baseline

from the gross area.)

F3 EXPAND Pressing the F3 key will expand the display.  The channel under

the cursor will try to stay near the middle of the graphic screen.

F4 CONTRACT Pressing the F4 key will contract the display.  The channel

under the cursor will try to stay near the middle of the graphic screen.

F5 IDENT Pressing the F5 key will cause the program to search through

the library, trying to find a match for the energy under the cursor.  At the

end of the search, the number of matches will be displayed near the bottom of

the screen.

F6 RESULT Pressing the F6 key will display the results of the most

recent search.  Examine one of the library entries by typing its number, or

return to the display by pressing RETURN.  If you type a valid number, then

you are presented with all of the isotope's data stored in the library.  Also

shown is the mass of the sample and geometry.  The detector efficiency and

amount of isotope in the sample are calculated and shown.  You now have the

option of printing this information on the printer, if desired, press Y,

otherwise press N or RETURN.  The isotope amount is calculated using one of

the following equations:

pCi/g = #           Net.area x  1 x 10#12#         #                     (Eq. 1)

Live x Det.eff x t.p x 3.7 x 10#10# x mass

Net.area is the net area of the spectral region, 1 x 10#12# is the

number of pico units, Live is the live time in seconds, Det.eff is the

detector efficiency, t.p is the transition probability, 3.7 x 10#10# is the

number of disintegrations per Curie, and mass is the sample mass in grams.

ppm = #        Net.area/Live x At.wt x 1 x 10#6#                   #     (Eq. 2)

Ab x t.p x Det.eff x flux x mass x c.s x Avog x irr x cool

Net.area is the net area of the spectral region, Live is the live time in

seconds, At.wt is the atomic weight in grams per mole, 1 x 10#6# is the number

of parts in one million, Ab is the isotopic abundance, t.p is the transition

probability, Det.eff is the detector efficiency, flux is the neutron flux in

neutrons per cm#2# per second, mass is sample mass in grams, c.s is the cross

section in barns, Avog is Avogadro's number, irr is the activation factor, and

cool is the decay factor:

irr = (1-(exp(-(ln(2)/HalfLife) x lenirr )))        (activation factor)

HalfLife and lenirr (length of irradiation) are in hours,

cool = (exp(-(ln(2)/HalfLife) x cooltime x 24 ))    (decay factor)

HalfLife is in hours and cooltime is in days.

The isotopic amounts for naturally occurring isotopes are calculated

using equation 1 (in programs ENRAD and ALPHA), while isotopes from thermal

neutron activated samples are calculated using equation 2 (in the program

GAMMA).

F8 SCALE Pressing the F8 key will allow changing the plot parameters,

such as connected dots and vertical scale.

F9 -MORE- Pressing the F9 key will present a second set of functions.

F1 ENERGY Pressing the F1 key will display the energy calibration factors

near the bottom of the screen.  These are the channel zero energy and the

energy per channel.  Both are in units of kev.

F2 TIME Pressing the F2 key will display the count times near the bottom

of the screen.  These are clock and live times in seconds.  The percent dead

time is calculated from these times.

F3 MEMORY Pressing the F3 key will display the available memory.

F4 HEADING Pressing the F4 key will display the heading information that

is stored along with the spectral file.

F5 DETEFF Pressing the F5 key will cause the program to plot the

efficiency curves for the various detector geometries that are currently

defined.

F6 SAVE Pressing the F6 key will cause the program to write the entire

spectral file back to disk.  This is useful if the spectrum is recalibrated

and then the user wishes to store the new calibration coefficients.

F7 CALIB Pressing the F7 key will allow the user to recalibrate the

spectrum.  The program will prompt the user with instructions printed near the

bottom of the screen.  First, move the cursor to a peak of known energy, press

the RETURN key, and enter the energy (in kev).  Then, move the cursor to a

second peak of known energy, press the RETURN key, and enter the energy (in

kev).  The program will calculate the channel zero energy and the energy per

channel and display them near the bottom of the screen.  These new

coefficients can be stored with the spectrum by pressing the F6 SAVE key.

F8 LOAD Pressing the F8 key will cause the program to prompt the user for

another file name.  This spectral file will be loaded and displayed.

F9 -MORE- Pressing the F9 key will present the first set of functions.

#Coefficients File

The file called COEFFS.TXT contains the detector geometry coefficients.

The format of the file is described as follows.  The first line of the file

contains the number of geometries in the file.  Each geometry is given:

A)  two coefficients, called a and b.  These coefficients are used to

determine the detector efficiency at any given energy by solving this

equation:

Det.Eff. = a * energy#b#

(**NOTE** This equation gives erroneous efficiencies for energies below

about 100 kev.)  These coefficients can be determined by counting a

radioactive standard (available from the National Bureau of Standards),

and plotting the ratio of detected activity to known activity versus

energy.  This data can be fitted to the above equation using a least

squares method.  This method is carried out for each detector geometry

used, and the data is put into the coefficients file.

B)  the title, or name, of the geometry.  This is usually a number, 1, 2,

etc., for each geometry used.

#Configuration File #

A configuration file can be created to avoid having to answer the plot

questions every time a spectral analysis program is run.  If this file is

present on the default drive, then the program will take answers to the plot

questions from this file.  The configuration file should contain two or three

pieces of information.  The first character should be Y or N, and is the

answer to the question "Connect dots? (Y/N)".  The second character is also

either Y or N, and is the answer to the question "Logarithmic plot? (Y/N)".

If the second character is Y (which means that log plot is selected), then

that is all the configuration file needs to contain.  If the second character

is N (which means that log plot is not selected), then a numeric value will

have to be provided.  To cause the spectral analysis program to always scale

the spectrum to the largest value, enter a zero here; otherwise, any non-zero

value can be entered.  The simplest method of creating a configuration file is

illustrated here: (user input is italicized; <CR> means to press the RETURN

key)

Comments

A>#copy con gamma.cfg <CR>#  copy from console to configuration file

#NN0<F6>  <CR>#            <F6> means press the F6 key

1 File(s) copied         response from DOS

#Spectrum Acquisition

The spectrum will have to be acquired from a multi-channel analyzer or

similar instrument, and then processed into the particular format required by

the analysis program.  The program called ACQUIRE3.BAS is a program written in

BASIC which was designed to operate with a model 5400 multi-channel analyzer

built by Norland/Ino-Tech, using a serial interface.  It will ask for some

heading information to be stored along with the spectrum, and then will

display some instructions before acquiring the spectrum.  This program is very

picky and probably will have to be modified to work with other equipment.  But

it should give a rough idea on how to write a program to acquire and process a

spectrum.

#Spectral File Format #

The spectral file must be in a particular format in order for the

programs GAMMA, ENRAD, and ALPHA to read them properly.  The program XRAY has

slightly different file format requirements.  In the following, the string

called "miscellaneous" will be printed on the top line of the graphic screen.

Usually the title of the spectrum is put here.

For the GAMMA, ENRAD, and ALPHA programs the format is as follows:

#Name of variable           variable type        description

irrdate                    string[80]      date of irradiation

irrtime                    string[80]      time of irradiation

colldate                   string[80]      date of collection

colltime                   string[80]      time of collection

misc                       string[80]      miscellaneous

flux                       real            neutron flux in units of

neutrons per cm#2# per second

mass                       real            sample mass in grams

geom                       integer         detector geometry

lenirr                     real            length of irradiation in hours

cooltime                   real            cooling time in days

chzen                      real            channel zero energy in kev

ench                       real            energy per channel in kev

numchan                    integer         number of channels in spectrum

clock                      real            clock time in seconds

live                       real            live time in seconds

counts                     real            counts in channels 2 through

numchan.  Channels 0 and 1 are

assumed to contain the clock and

live times, respectively.

<EOF>                                      end of file.

The following file format applies to the program XRAY.COM:

#Name of variable           variable type        description

misc                       string[80]      miscellaneous

startang                   real            starting angle of scan

endang                     real            ending angle of scan

angincr                    real            angle increment

CTime                      integer         count time in seconds

...

numchan                    integer         number of channels

counts                     real            counts in channels 1 through

numchan

<EOF>                                      end of file.

#Isotope Libraries

Three isotope libraries are included.  They are NTHERM2.TXT, ENRAD2.TXT,

and NATALP2.TXT.  They include information such as atomic number and atomic

mass of isotopes, cross section and abundance of parents, and emission

energies and transition probabilities.  Up to 24 energies and transition

probabilities are included with each isotope.  This data has been collected

from General Electric Nuclide Chart.

The library NTHERM2.TXT contains isotopes which are produced by thermal

neutron irradiation.

The library ENRAD2.TXT contains isotopes which are naturally occurring

gamma emitters.  The isotopes are products of the uranium and thorium decay

chains.

The library NATALP2.TXT contains isotopes which are naturally occurring

alpha particle emitters.